## Calibration
Authorized User: |  |  |  | Instrument Type: |  | LSC
Building: |  |  |  | Manufacturer/Model:
Room: |  |  |  | Serial No.:
Calibration Date: |  | 2024-02-02
Radionuclide | Serial No. | Instrument Setting (Window) | Channel/Region | Sample Count Time (min) | Background Count Time (min) | Background (CPM) | Gross Sample (CPM) | Net Sample (CPM) | Standard Assay Activity (DPM) | Standard Assay Reference Date | Half-Life (days) | Decayed Activity (DPM) | Efficiency | Calibration Factor | MDA (DPM)
P-32/Si-32 |  | 256-2000 | C | 1 | 1 | 12 | 59822 | 59810 | 75180 | 2019-04-30 | 57305 | 73615.1408 | 0.81247 | 1.2308 | 23.5303
C-14 |  | 18.5-256 | B | 1 | 1 | 8 | 64845 | 64837 | 99900 | 2018-10-08 | 2080500 | 99835.3519 | 0.64944 | 1.5398 | 24.883
H-3 |  | 2-18.6 | A | 1 | 1 | 15 | 82511 | 82496 | 213120 | 2018-10-08 | 4496.8 | 157962.3734 | 0.52225 | 1.9148 | 40.249
Swipe Protocol #:
Calibration performed by:
Unknown Efficiency: |  | 52.23%
How to use this spreadsheet:
Step 1: | All highlighted fields are where you enter your data. Example data has been entered for demonstration purposes and will need to be updated with your data.
Step 2: | RSS recommends counting your samples for 1 minute. Most of the time, LSC standards sets come with a background, C-14, and H-3 standard. If you have more sources to count, you can adjust the above template as you need.
Step 3: | Make sure to enter each standard's Assay Date and Assay Activity so the spreadsheet can decay correct it to today's date.
Step 4: | The Calibration Date above will auto populate with the day you open the spreadsheet.
Step 5: | If you have an older LSC such as a Beckman, you will need to make sure your swipe protocol has a calibration factor set to 1 for each isotope before you collect data. This will ensure that your results will be in CPM. See the red circle below for the calibration factors.
Step 6: | Once you collect data, enter the background cpm for each region. Data from the image below has been entered as an example
Step 7: | Enter the data for each standard. Only enter the cpm that corresponds to the isotope and region of interest. See the highlighted portions in the below example, data has been entered above as an example.
Step 8: | Make sure the efficiency above looks correct. Typical efficiencies for H-3 will be 50-65%, C-14 will be 65-75%, and P-32 will be 70-90% for older LSCs.
Step 9: | Make sure your MDAs are <200 dpm which is the trigger level for removable contamination.
Step 10: | If the data looks good, then for older LSCs such as Beckmans, update the swipe protocol with the new calibration factors.
Step 11: | Go to sheet "Print this out", print out the calibration information, and place it on your LSC. Everything on the calibration sheet will auto populate from this sheet.
Step 12: | Send a final copy to RSS at rad@missouri.edu
<--Calibration Factors set to 1, so data is in cpm.
B1: Background
B2: C-14
B3: H-3
B4: P-32/Si-32

## Print this out
Calibration Sticker
Radionuclide | Efficiency | MDA (DPM)
P-32/Si-32 | 81.25% | 23.5303
C-14 | 64.94% | 24.883
H-3 | 52.23% | 40.249
Unknown | 52.23% | 23.5303
Swipe Protocol #: | 0
Calibrated by: | 0
Calibration Date: | 2024-02-02